Integrated Sealed EO Vapour Test Cell

Dimensioning basis — current handoff release, 2026-09-03

Status: authoritative engineering concept; fabrication release is blocked by the open evidence gates.

| NET-VOLUME TARGET / SIDE | MEMBRANE REGION | LOCAL GAPS |
| --- | --- | --- |
| 500.0 +/- 5.0 mL | 50 x 50 mm (25 cm2) | 1 / 2 / 4 mm |

# Selected physical system

Two equal sealed gas chambers, A and B, share the active membrane boundary. Each assembled chamber shall be calibrated to 500.0 +/- 5.0 mL net free-gas volume at 32 °C, 101325 Pa and 50% initial RH. The membrane coupon is 80 x 80 mm, the active region is 50 x 50 mm, the clear chamber face is 100 x 100 mm, and the outer plate envelope is 110 x 110 mm. The volume is a governing design target; traceable assembled-cell calibration remains open.

Figure 1. Dimensioned exterior and section engineering concept.

# Volume derived from target drift and switching time

The declared target flux is 208 g m-2 h-1 across 25 cm2, giving 520.0 mg/h (0.144444 mg/s). A 5.0 s polarity segment transfers 0.722222 mg. Using V_required = transferred mass / (saturated vapour density x 0.05), the minimum net free-gas volume is 427.6 mL per side. The selected 500.0 mL target gives a ratio of 1.169 and a predicted change of 4.276 RH percentage points in either chamber at the first switching point.

# Local, bulk, net and gross volume

| Gap | Local membrane region | Bulk free gas | Net-volume target |
| --- | --- | --- | --- |
| 1 mm | 2.5 mL | 497.5 mL | 500.0 mL |
| 2 mm | 5.0 mL | 495.0 mL | 500.0 mL |
| 4 mm | 10.0 mL | 490.0 mL | 500.0 mL |

The local membrane-region volume is included in the net-volume target. Gross empty-cavity volume is released from gross = calibrated net + installed component displacement after the internal hardware inventory is measured. The rendering does not define any dimension stored as null in the configuration.

# Ideal ABBA transient

| Time | RH A | RH B | Transfer A to B |
| --- | --- | --- | --- |
| 0 s | 50.000% | 50.000% | +0.000000 mg |
| 5 s | 45.724% | 54.276% | +0.722222 mg |
| 10 s | 50.000% | 50.000% | +0.000000 mg |
| 15 s | 54.276% | 45.724% | -0.722222 mg |
| 20 s | 50.000% | 50.000% | +0.000000 mg |

The analytic +1, -1, -1, +1 sequence conserves total water inventory, gives equal and opposite chamber changes, and returns both chambers to the initial state at 20 s. The gas-only pressure diagnostic reaches 406.9 Pa at an extremum. Pressure management remains open pending verification against the 5.0 Pa target.

# Engineering register for fabrication release

Fabrication release ready: false. Every required register item must be closed with cited evidence before this may be set to true. The gates explicitly cover walls and divider, ports and feedthroughs, fan, baffles, fixture mounting, fasteners, datums, tolerances, gap uniformity, pressure management, mixing response and net-volume calibration.

| Item | Status | Closure evidence |
| --- | --- | --- |
| Net Free Gas Volume Target | Defined | Configuration-controlled 500.0 +/- 5.0 mL net free-gas acceptance target per assembled side |
| Gross Empty Cavity Volume And Depth | Open | Released cavity geometry and dimensional inspection |
| Installed Component Displacement | Open | Displacement inventory for fan, probes, baffles and ports |
| End Wall And Central Divider Structural Sizing | Open | Selected materials and thicknesses supported by structural calculation and loaded-deflection inspection |
| Fluidic And Electrical Port Interfaces | Open | Released interface specification covering threads, seals and feedthroughs plus an assembled leak test |
| Internal Fan Selection And Mounting | Open | Selected fan, envelope and mount with measured heat load and chamber displacement |
| Baffle Geometry And Mounting | Open | Released baffle CAD and mounting details supported by the chamber mixing test |
| Differential Pressure Management | Open | Transient pressure test across the full polarity sequence |
| Fixture Mounting Fastener Pattern And Preload | Open | Released external mounting interface, bolt-load calculation and torque map |
| Seal Geometry And Compression | Open | Seal supplier limits and compression inspection |
| Drawing Datums And Dimensional Tolerances | Open | Released datum and tolerance scheme with completed dimensional inspection |
| Loaded Flatness And Gap Uniformity | Open | Mapped 1, 2 and 4 mm gaps under operating preload |
| Internal Mixing Time | Open | Humidity step-response test in each chamber |
| Net Volume Calibration | Open | Traceable measured 500.0 +/- 5.0 mL free-gas volume calibration for each assembled side |

# Reproduction and verification

Run scripts/design_eo_vapor_test_cell.py to regenerate the volume basis, 0.1 s transient, engineering register, audit and source manifest. Run pytest -q tests/test_eo_vapor_test_cell_design.py for the formula, conservation, symmetry and packaged-output checks.