EO Vapour Test Cell

Selected design basis for net-transient measurement

Engineering working basis · calculation-backed concept · not released for fabrication

| ACTIVE MEMBRANE | NET VOLUME / SIDE | WORKING SEGMENT |
| --- | --- | --- |
| 50 x 50 mm (25 cm²) | 500 ± 5 mL | 12 s |

Figure 1. Two-dimensional functional blueprint of the complete test setup. Values and dimensions are kept in the calculation basis; this is not released fabrication CAD.

# Selected values and reasons

| Quantity | Selected value | Why it was selected |
| --- | --- | --- |
| Active membrane | 50 x 50 mm = 25 cm² | Largest existing aperture; strong signal and lower edge-to-area ratio. |
| Membrane coupon | 80 x 80 mm | Provides land for double seals, guard groove and clamping. |
| Local gas gap | 2 mm per side | Matches the local transport model; active-gap volume is only 5.0 mL per side. |
| Calibrated net gas volume | 500 ± 5 mL per side | Retains the illustrated 0.50 L cell and slows the transient for the present instrumentation. |
| Bulk plenum | 165 x 165 x 20 mm | Provides gross volume without changing the 25 cm² active membrane. |
| Polarity sequence | 12 s each; +I, −I, −I, +I | Measures net transfer by polarity while suppressing fixed and linear drift. |

# Calculation basis

The calculation is conditional on 208 g m⁻² h⁻¹ at 32 °C, 101.325 kPa and 50% initial RH. The flux is a test-sizing target, not a measured material property.

ṁ = J A = 208 × 0.0025 = 520 mg/h = 0.144444 mg/s

Δm = ṁ Δt = 0.144444 × 12 = 1.733333 mg

ΔRH = Δm / (ρsat Vnet) = 10.261 percentage points per chamber

After the first positive segment, the ideal well-mixed endpoints are 39.739% and 60.261% RH. The differential RH signal is 20.523 percentage points.

# Volume closure

544.5 mL gross − 22.6 mL components − 21.9 mL trim = 500.0 mL net

The gross 165 x 165 x 20 mm plenum is 544.5 mL. After the conservative 22.6 mL component bound, 521.9 mL remains. Removable inserts must cover 21.9–44.5 mL until actual solid displacement is measured.

# Instrument and pressure implications

The SHT45 step-response fraction over 12 s is 95.0%. This does not mean exact ramp tracking: a simple first-order model indicates only 68.3% of a linear endpoint change. Use a faster calibrated reference if the transient shape is required. The rigid-cell pressure diagnostic is 976 Pa, compared with the 5 Pa target. The ideal pressure-management screen therefore requires about ±2.40 mL travel per side and 4.79 mL full stroke.

# Interpretation boundary

Report: integrated net water moved through the membrane during each segment under evolving RH conditions. Do not claim: constant-boundary intrinsic EO flux or sustained dehumidification from the closed symmetric sequence alone. One four-segment block lasts 48 s and returns the ideal inventory to its initial state.

Fabrication remains conditional on measured mixing, leakage, fan thermal bias, differential pressure, component displacement, seal performance and assembled net volume.

# Error-controlled construction

The cell geometry is now tied to an explicit error budget. Calculation is used for stiffness and inventory scaling; physical blanks and scaling tests control effects that cannot be credibly removed by geometry alone.

| Control | Selected design / acceptance criterion |
| --- | --- |
| Structural wall | 12 mm EN AW-6082-T6 aluminium; 1.33 µm pressure-only proof deflection. |
| Insulating plate | 20 mm Unfilled PEEK; 4.51 µm pressure-only proof deflection. |
| Loaded local gap | 2.00 ± 0.02 mm with hard stops and a loaded gap map. |
| Wetted surfaces | Smooth removable non-porous metal liners; exposed PEEK restricted to the insulating cartridge land. |
| Edge effect | Primary 50 mm aperture plus 30/40/50 mm area–perimeter regression. A 1 mm band is already 7.84% of active area. |
| Wall-storage blank | ≤ 0.01733 mg per segment, equivalent to 0.256 nm reversible water film. |
| Sensors / mixing | 2 T/RH locations per chamber; relative calibration ≤ 0.10 RH point and spatial disagreement ≤ 0.15 RH point. |
| State / leakage | Chamber ΔT ≤ 0.05 K; volume uncertainty ≤ 1.0 mL; cross-leak screen ≤ 6.25 mL/min at peak RH difference. |

# Acceptance logic

The planning root-sum-square allocation is 3.36% of the corrected segment signal. This is not demonstrated accuracy. Accept data only after individual volume calibration, pressure-decay and guard-leak testing, an impermeable-coupon wall/thermal blank, a zero-current diffusion blank, three-aperture edge regression, spatial mixing verification and post-run drift checks.

Wall thickness cannot eliminate membrane-edge transport or reversible water storage. These are controlled by the liner, guard, blanks and aperture regression.